## Slide 1: Asbestos measurement techniques, monitoring ambient and indoor air using SEM, methods and results of a first international PT scheme. Authors/presenter: R. Koenig, APC GmbH, Eschborn, Germany; L. Davies, HSL, Buxton, Great Britain; N. Hoefert, VDI, Düsseldorf, Germany; Jan Tempelman, TNO, Utrecht, Netherlands

## Slide 2: ISO 14966 / VDI 3492
The above mentioned guidelines (with some modifications) using Scanning Electron Microscope (SEM) based measurement methods are widely used
for the determination
of the asbestos fibre concentration in indoor and ambient air in Germany and to some extent in some of its neighboured European countries

## Slide 3: PT Schemes
The results of measurements of the asbestos fibre concentration are used to make decisions, which may result in financial and economical consequences, in addition to the health effects of workers and others.

## Slide 4: PT Schemes
APC GmbH 2011
As a result it is critical, that laboratories operating in this field are giving comparable results. It is therefore surprising, that -with the exception of two round robins in Germany during the 1990’s - no regular PT scheme is available for this technique.

## Slide 5: PT Schemes
APC GmbH 2011
The main reason for this is the technical difficulty of producing large numbers of filters each with equivalent fibre densities, in a suitable concentration range and at a reasonable cost.

## Slide 6: PT Schemes
APC GmbH 2011
To address this problem a small consortium of interested parties was formed (namely APC and VDI in Germany, TNO in the Netherlands and HSL in the UK), with the remit of developing a suitable SEM PT scheme.

## Slide 7: ISO 14966 PT Scheme
APC GmbH 2011
From the beginning we were aiming for an international acceptable scheme, therefore the basis was ISO 14966,
being also very similar to VDI 3492 and other SEM methods

## Slide 8: ISO 14966 PT Scheme
APC GmbH 2011
The intention was to produce a larger number of equally loaded filter cassettes
with fibre concentrations up to 50 fibres / mm² filter area.
The sample set for the participants shall include at least three filter cassettes having different qualities both in fibre number and asbestos type on the filter.

## Slide 9: ISO 14966 PT Scheme (steps included)
APC GmbH 2011

## Slide 10: ISO 14966 PT Scheme (production of loaded filter cassettes)
APC GmbH 2011
Sample production is achieved by HSL using a combination of a modified dustiness tester to generate airborne fibre concentrations and a multi-port sampling device – the “Sputnik”.
A combination of commercial asbestos containing materials and specially prepared spiked materials was the source for generating the dust for the filter loading

## Slide 11: ISO 14966 PT Scheme (production of loaded filter cassettes)
APC GmbH 2011
Using a critical orifice and a sampling flow rate of around 1.8 litres per minute, the “sputnik” has the capability to collect up to 114 equally loaded filters (mounted in cassettes), in one batch.
In the first pilot PT scheme two differently loaded filters were used together with an ambient air filter generated by parallel sampling at the same location.

## Slide 12: ISO 14966 PT Scheme (production of loaded filter cassettes)
APC GmbH 2011

## Slide 13
APC GmbH 2011

## Slide 14: ISO 14966 PT Scheme Additional Information Package
APC GmbH 2011
An additional information package
was sent together with the samples, to ensure with other things, that the counting rules are well understood. Most people use WHO counting rules in their every day business, which are different in some cases from ISO counting rules

## Slide 15: Pilot PT Scheme ISO 14966
APC GmbH 2011
During this first trial the number of participants was limited to 15 laboratories, representing five European countries (Austria, Germany, Great Britain, Netherlands and Switzerland).

## Slide 16
APC GmbH 2011
Pilot PT Scheme ISO 14966 Results Fibre numbers (L > 5 µm) |  |  |  |  |  |  | 
Laboratory | Ambient air |  | R 24 |  | R 25 |  | Evaluated filter area mm2
 | Other In. | Asbestos | Other In | Asbestos | Other In. | Asbestos | 
A | 22 | 0 | 1 | 6 | 22 | 30 | 1,05
B | 0 | 0 | 0 | 10 | 0 | 32 | 1,04
C | 8 | 0 | 3 | 6 | 2 | 31 | 1,03
D | 5 | 0 | 0 | 7 | 2 | 22 | 1,01
E | 3 | 0 | 0 | 11 | 1 | 42 | 1
F | 0 | 0 | 4 | 13 | 11 | 18 | 1
G | 0 | 0 | 0 | 10 | 0 | 18 | 1
H | 1 | 0 | 1 | 3 | 0 | 16 | 1,01
I | 7 | 1 | 0 | 9 | 5 | 38 | 1
K | 0 | 0 | 0 | 5,3 | 0 | 12,9 | 1,7
L | 0 | 0 | 0 | 5 | 0 | 24 | 1,01
N | 0 | 0 | 0 | 8 | 1 | 14 | 1,01
O | 0 | 0 | 0,5 | 2,7 | 0 | 7,7 | 1,83
P | 11 | 0 | 3 | 8 | 1 | 36 | 1
Q | 1 | 0 | 1 | 9 | 0 | 26 | 1
green numbers: calculated for 1 mm² evaluated filter area |  |  |  |  |  |  | 

## Slide 17
APC GmbH 2011
Chart: Distribution of fibre diameter (asbestos) as detected by the participants
| 0.1 | 0.2 | 0.30000000000000021 | 0.5 | 0.8 | 1.5 | 2.5
%/µm | 0.068532 | 0.37122 | 0.31125 | 0.18561 | 0.054255 | 0.0085665 | 0.0005711

## Slide 18: Pilot PT Scheme ISO 14966
APC GmbH 2011
Asbestos varieties in the sample:
50 % crocidolite
45% amosite
5% chrysotile

## Slide 19: Pilot PT Scheme ISO 14966 Results
APC GmbH 2011
Observed Deviations:
In a simple estimation the deviation SR resulting from the evaluation by the participants extending the unavoidable poissonian random variation SP can be calculated:
where SN is the standard deviation of the counting results for one filter type

## Slide 20: Pilot PT Scheme ISO 14966 Results
APC GmbH 2011
Observed deviations in between-laboratory comparison:
For the ambient air filters no deviation extending the poissonian uncertainty detectable
For the filter type R 24 the variation SR, which results from the between-laboratory comparison of the fibre counts is: appr. 10% relative
For R 25 the variation SR is appr. 30% relative

## Slide 21: Pilot PT Scheme ISO 14966 Results
APC GmbH 2011
Detected reasons for the observed variations:
Laboratories with a tendency to be lower than average (asbestos) show two types of discernable deviations:
Smallest detected fibre diameter is 0,3 µm or greater according to their fibre counting form.
or
they record a larger amount of fibres with “no spectra” = unidentified objects

## Slide 22: Pilot PT Scheme ISO 14966 Results
APC GmbH 2011
Reasons for the observed variations:
Optimization of EDXA instruments
required:
Greater channel width (20 eV or even more) gives
maximum P/B in shortest possible measurement time and is especially important for the analysis of very thin fibres, to minimize statistical uncertainties in the spectra and to reduce the adverse effects of beam drift during analysis.

## Slide 23: ISO 14966
APC GmbH 2011
Optimization of EDXA instruments required:

## Slide 24
APC GmbH 2011
Optimization of EDXA instruments

## Slide 25
APC GmbH 2011
Optimization of EDXA instruments

## Slide 26: Pilot PT Scheme ISO 14966 Summary
APC GmbH 2011
The pilot PT scheme was running as scheduled
For a first PT scheme including both preparation
and analysis of the filter samples
the results can be regarded as good, especially in the range of the decision limits (clearance levels as an example) in Germany and some of its neighbored countries.
Based on the experience HSL (Health and Safety Laboratory) will offer a full scale PT scheme starting out later this year.

## Slide 27
Thank you for your attention !